CT6 Increased Accuracy Calibration/Test System
This approach to calibration improves on the previous approach providing the ability to calibrate the voltage and current more accurately than the previous system and at ~ 50A load (although only using a 2 kW fan heater). This ensures calibration is performed in a linear section of the 100A CT performance curve. A WiFi network must be available to allow connectivity between the PC, Shelly 1PM Plus unit and the CT6 unit under test (UUT).
Micro USB
connector
Modified USB Cable.
The display
Should be connected to this connector
PC running Ubuntu 24.04
This AC – AC PSU (transformer) must provide
9 to 16 (AC) volts to the CT6 unit.
! The calibration process is only valid for the PSU type tested.
AC
PSU
SCT013
Current
Transformer
Live wire is looped through the
SCT013 current transformer
so that it passes through it 6
times.
Calibrated
Shelly 1PM Plus
230V AC
Mains
Supply
L
2 kW
Electric
Heater
N
E
L
N
A 2 kW fan heater
is used as the load.
E
Note
The voltage and current readings from the
Shelly 1PM Plus unit should be calibrated using accurate voltage and AC current meters. See the pypi, test_equipment, python module for details of this.
Note !!!
Ensure all wiring is safe as mains AC voltage is dangerous.